## Dash Board
[17 empty] | Equipment Photo: |  | Photo 1
Name/Company : | SAVVICE CORPORATION |  |  |  |  | Place of Calibration : |  | Force Standards Laboratory, NML ITDI |  |  |  |  | Machine Indication
Address Line 1 : | 3rd Floor BMWC Bldg., Cagayan Valley Road | [11 empty] | 0 | 0
Address Line 2 : | Sta. Rita, Guiginto, Bulacan |  |  |  |  | Person in Charge : | [6 empty] | 1st | 3000
Requesting Official : | Ramil F. De Luna |  |  |  |  | Analyst/Calibrator : |  | Ahdrian Camilo C. Gernale |  |  |  |  | 2nd | 6000
Position : | Technician | [6 empty] | Radley F. Manalo |  |  |  |  | 3rd | 9000
[13 empty] | 4th | 12000
[13 empty] | 5th | 15000
6th
[13 empty] | 7th | [5 empty] | Photo 2
8th
Date of Calibration : | 2026-04-22 |  |  |  |  | Calibration Mode |  | Compression |  |  |  |  | 9th
Req. Reference No. : | 03-2026-FORC-0087 |  |  |  |  | Capacity : |  | 15000 | kgf |  |  |  | 10th
Calibration Item : | Axle Weighing Scale |  |  |  |  | Range : |  | 0 |  | 15000 |  |  | 0 | 0
Calibration Type : | [5 empty] | Increment | : | 50
Load Cell Type : | No |  |  |  |  | Resolution : |  | 50
Force Transducer Info
Make / Model No. : | LS 630 W. PAD
Serial No. / ID No. : | 0122YL21002
Indicator Info
Make / Model No. : | INTERCOMP PT20
Serial No. / ID No. : | 0307YM23003
Special Instructions & Remarks :
please input above the preparations and other pertinent details for the calibrators to remember before starting
Calibrator Tools :
TSR # | Date of Calibration | In-Charge/ Supervisor |  | Location | Standards Used | Calibrator/ Analyst | Test Points | [9 empty] | Equivalent Force | [9 empty] | U | Class | Remarks
[7 empty] | 1st | 2nd | 3rd | 4th | 5th | 6th | 7th | 8th | 9th | 10th | 1st | 2nd | 3rd | 4th | 5th | 6th | 7th | 8th | 9th | 10th
Note on D7: This will be displayed in the certificate

## Software (hidden)
FORCE MEASUREMENT DATA
Ver. 2 Rev.0 Mar 12, 2019
revised & developed by: gearnhale@yahoo.com according to ISO
original developer: manuelmruiz@yahoo.com
Date of Calibration |  | Calibration Mode
Request Reference No. |  | Capacity
Calibration Item |  | Range
Make / Model No. |  | Increment/Graduation
Serial No. / ID No. |  | Resolution
Guide on how to use this data entry system: |  |  | Environmental Conditions:
1. Select the radio button of the machine indicator unit (kgf, lbf, N, kN). |  |  |  | Before | After
2. After capturing the digital amplifier's reading, change the machine indication accordingly. |  |  | Temp: | 25 | 25
3. Upon completion of each loading set, press the corresponding PRINT button. Make sure you have |  |  | Humidity: | 1900-02-01 | 1900-02-01
change first the machine indication reading upon capturing.
Data Logger |  |  | Calibration Worksheet
Machine Indicator Readability
Zero | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
| Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0 |  | Point No. |  | kN |  |  | raw | zero corrected
|  |  | Zero | 0
1st Test point | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 |  | 1st | 20
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | 2nd | 40
| Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0 |  | 3rd | 60
|  |  | 4th | 80
2nd Test point | 1.1 1.1-DMP40 -0.242682 mV/V abs 13-Mar-19 13:06:55 |  | 5th | 100
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | 6th
| Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0 |  | 7th
8th
3rd Test point | 1.1 1.1-DMP40 -0.403444 mV/V abs 13-Mar-19 13:07:17 |  | 9th
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | 10th
| Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0 |  | Zero | 0
Force Transducer Information
4th Test point | 1.1 1.1-DMP40 -0.563466 mV/V abs 13-Mar-19 13:08:12 | [8 empty] | 0.0098067 | kgf
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | Load Cell | HBM/C3H3 |  |  |  | A
| Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0 |  | Capacity |  |  |  |  | B
|  |  | Serial No. |  |  |  |  | C
5th Test point | 1.1 1.1-DMP40 -0.723624 mV/V abs 13-Mar-19 13:09:02 |  | Cal. Cert. No. |  |  |  |  | U, %
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | Cal Date |  |  |  |  | k | 2
Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0
6th Test point |  |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
|  |  | Point No. | kgf | kN |  |  | raw | zero corrected
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0
|  |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971
7th Test point |  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475
Pre-Load 2
|  |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
8th Test point |  |  | Point No. | kgf | kN |  |  | raw | zero corrected
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0
|  |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475
9th Test point
Pre-Load 3
|  |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
|  |  | Point No. | kgf | kN |  |  | raw | zero corrected
10th Test point |  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0
|  |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475
Zero | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 |  | Series 1
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
| Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0 |  | Point No. | kgf | kN |  |  | raw | zero corrected
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0 |  | Error Estimation, in %
|  |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971
|  |  | 2nd | 40 | 0.39227 | 1.1 1.1-DMP40 -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.24268 | -0.32077 | -1.5727
|  |  | 3rd | 60 | 0.5884 | 1.1 1.1-DMP40 -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.40344 | -0.48153 | -2.361
|  |  | 4th | 80 | 0.78453 | 1.1 1.1-DMP40 -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.56347 | -0.64155 | -3.1457
|  |  | 5th | 100 | 0.98067 | 1.1 1.1-DMP40 -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.72362 | -0.80171 | -3.9311
|  |  | 6th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 7th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 8th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 9th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 10th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475
Series 2
|  |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
|  |  | Point No. | kgf | kN |  |  | raw | zero corrected
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0 |  | Error Estimation, in %
|  |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971
|  |  | 2nd | 40 | 0.39227 | 1.1 1.1-DMP40 -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.24268 | -0.32077 | -1.5727
|  |  | 3rd | 60 | 0.5884 | 1.1 1.1-DMP40 -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.40344 | -0.48153 | -2.361
|  |  | 4th | 80 | 0.78453 | 1.1 1.1-DMP40 -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.56347 | -0.64155 | -3.1457
|  |  | 5th | 100 | 0.98067 | 1.1 1.1-DMP40 -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.72362 | -0.80171 | -3.9311
|  |  | 6th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 7th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 8th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 9th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 10th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475
Series 3
|  |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
|  |  | Point No. | kgf | kN |  |  | raw | zero corrected
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0 |  | Error Estimation, in %
|  |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971
|  |  | 2nd | 40 | 0.39227 | 1.1 1.1-DMP40 -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.24268 | -0.32077 | -1.5727
|  |  | 3rd | 60 | 0.5884 | 1.1 1.1-DMP40 -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.40344 | -0.48153 | -2.361
|  |  | 4th | 80 | 0.78453 | 1.1 1.1-DMP40 -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.56347 | -0.64155 | -3.1457
|  |  | 5th | 100 | 0.98067 | 1.1 1.1-DMP40 -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.72362 | -0.80171 | -3.9311
|  |  | 6th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 7th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 8th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 9th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 10th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475

## Calibration Data (hidden)
FORCE MEASUREMENT DATA
Ver. 2 Rev.0 Mar 12, 2019
revised & developed by: gearnhale@yahoo.com according to ISO
original developer: manuelmruiz@yahoo.com
Certificate No.
Client Name:
Calibration Date:
Client Address: | [8 empty] | Calibration Mode:
Calibration Item: | [8 empty] | Machine Capacity:
Model Code/No: | [8 empty] | Calibration Range:
Serial No/ ID No. | [8 empty] | Increment
Remarks: | [8 empty] | Resolution
1. Preloading
|  | Indicated Force | Load 1st (0˚) | Indicated Force | Load 2nd (120˚) | Indicated Force | Load 3rd (240˚) |  | Reference Standard Used: |  |  |  |  | Temp.: | Humidity:
a'
[11 empty] | b' | [14 empty] | Relative Expanded Uncertainty |  |  | Relative standard uncertainty of force measuring instrument wstd
c'
[18 empty] | RAW DATA CALCULATION | [8 empty] | Relative uncertainty by reproducibility (wrep,i) | Relative uncertainty due to resolution (wres,i) | Calibration uncertainty (wstdc) | Uncertainty due to temperature difference (wtemp') | Uncertainty due to sensitivity drift (wdrift) | Relative standard uncertainty (wstd) | Relative combined uncertainty (wci) | Effective degree of freedom | coverage factor (k) | Relative expanded uncertainty (Wi) | Relative accu error qi | Relative rep error bi | Relative zero error f0 | Relative accuracy error qi | Relative repeatability error bi | Relative zero error f0 | Lower Limit of calibration range | Relative uncertainty of force measuring instrument 상대불확도 | Class
| 2. Linear Interpolation (of reference force deflection) | [6 empty] | 3. Estimation of reference force | [9 empty] | Net Value (dij) & Mean (di) Calculation
|  | Force (kN) | Load Value | [5 empty] | Force (kN) | Load Value (kN) | [5 empty] | Machine Indication |  | Indicated Force | Load 1st (0˚) | Indicated Force | Load 2nd (120˚) | Indicated Force | Load 3rd (240˚) | Mean Value | Applied Force | Resolution testing machine
|  |  | 1st (0˚) | 2nd (120˚) | 3rd (240˚) | Mean Value |  |  |  | 1st (0˚) | 2nd (120˚) | 3rd (240˚) | Mean Value |  |  | Indicated Force | Equivalent Force (kN)
|  | 0 | 0 | 0 | 0 |  |  |  |  | -0.004913 | 0 | -0.004913 | 0 | [13 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | - | - | 0 | 0 | 0 | 0 | - | - | - | - | - | -
[26 empty] | Sensitivity on temp | [8 empty] | 2
2
[26 empty] | 0.003 | [8 empty] | 2
[26 empty] | Temp Diff. | [8 empty] | 2
2
[26 empty] | Drift Sens. | [8 empty] | 2
[26 empty] | 0.05 | [8 empty] | 2
2
2
2
|  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  |  | 0
[16 empty] | Max Zero | =

## Load Cell db (hidden)
Description | Capacity | Serial Number | Calibration Certificate | Calibration Date | a' (Compression) | b' (Compression) | c' (Compression) | U (Compression) | a' (Tension) | b' (Tension) | c' (Tension) | U (Tension) | Next Cal Date
HBM/C6 | 2000 kN | F42442 | 1800-00040-002 | 43153 | 978.3389 | 35.0482 | -7.9402 | 1.375 | -- | -- | -- | --
HBM/TCBS | 500 kN | 50036 | 1800-00040-003 | 43161 | -247.8101 | 0.47744 | -0.010141 | 0.07 | -- | -- | -- | --
HBM/C3H3 | 10 kN | F93007 | 1800-00040-004 | 43161 | 4.9026 | -0.00093075 | -6.3551e-05 | 0.011 | -4.9023 | 0.00041063 | -0.00016826 | 0.012
HBM/Z3H3 | 5 kN | F73724 | 1800-00040-005 | 43161 | 2.4495 | 0.00032711 | -0.00017387 | 0.017 | -2.4504 | -0.00031491 | -9.9356e-05 | 0.019
HBM/Z3OA | 2 kN | 151730166 | 1600-02019-001 | 42538 | 1000.037 | -0.091009 | -0.0017968 | 0.024 | -- | -- | -- | --
HBM/Z4A (20kN) | 20 kN | 161430192 | 1600-02019-003 | 42545 | 10.0026 | 0.00072857 | -0.00058347 | 0.019 | -10.0022 | 0.0018933 | 0.00041544 | 0.012
HBM/Z4A (100kN) | 100 kN | 160730020 | 1600-02019-002 | 42580 | 50.0339 | -0.014762 | -0.0022079 | 0.041 | -50.0346 | 0.0062677 | 0.0026657 | 0.053
WBRC/R200K | 200 kN | WB-1210-010 | 1600-02019-004 | 42576 | 90.7829 | -0.032822 | -0.0048756 | 0.081 | -- | -- | -- | --
WBRC-R2M | 2 MN | WB-121005 | 1800-00040-001 | 43153 | 1379.299 | 24.7927 | 18.2805 | 0.078 | -- | -- | -- | --
Shimadzu/ UH-F1000kNX | 1000 kN | 1251056K0094 | 11-2020-FORC-0116 | 44138 |  |  |  | 0.23 | -- | -- | -- | 0.1

## Data
CALIBRATION DATA SHEET
Customer Name: |  | ACME
Address Line 1: |  | 123 Road
Address Line 2:
Date of Calibration: | [6 empty] | Mode: | Compression
Request Ref. No.: |  | TEST |  |  |  |  | Capacity:
Calibration Item: | [6 empty] | Range:
Load Cell
Make / Model No.: | [6 empty] | Increment:
Serial No. / ID No.: | [6 empty] | Resolution:
Indicator
Make / Model No.:
Serial No. / ID No.:
Indicated force | 1st (0˚) | Indicated force | 2nd (120˚) | Indicated force | 3rd (240˚) |  | Conversion Unit:
| 0 | [7 empty] | kN
15000 | 14950
0
Calibration Points | Expected Machine Indication | Indicated force | 1st | Indicated force | 2nd | Indicated force | 3rd | Mean Force | Mean deflection
0 | 10 | 10 | 0 | 10 | 0 | 10 | 0
1st | 20 | 20 | 0 | 20 | 0 | 20 | 0
2nd |  |  | 6000 |  | 6000 |  | 6000
3rd |  |  | 9000 |  | 9000 |  | 9000
4th |  |  | 12000 |  | 12000 |  | 12000
5th |  |  | 14950 |  | 14950 |  | 15000
6th
7th
8th
9th
10th
0 |  |  | 0 |  | 0 |  | 0
Max Zero =
3.) Force Transducer Information | [6 empty] | 4.) Environmental Conditions:
Thermo-hygrometer indication
Load Cell |  |  | A |  |  |  |  | Before | After
Capacity |  |  | B |  |  |  | Temp:
Serial No. |  |  | C |  |  |  | Humidity:
Cal. Cert. No. |  |  | U, %
Cal Date |  |  | k
Calibrated by: | [5 empty] | Checked by:
RFM |  |  | ACCG
date:
5.) Net value (dij) & Mean Value (di) Calculation, units in kN
Calibration Points | Force (kN) | 1st (0˚) | Force (kN) | 2nd (120˚) | Force (kN) | 3rd (240˚) | Mean Force (kN) | Mean deflection
0 |  | 0 |  | 0 |  | 0
1st
2nd
3rd
4th
5th
6th
7th
8th
9th
10th
0
6.) Linear Interpolation | [5 empty] | 7. Estimation of Reference Force
Force (kN) | Load Value |  |  |  |  | Force (kN) | Load Value
| 1st (0˚) | 2nd (120˚) | 3rd (240˚) | Mean Value |  |  | 1st (0˚) | 2nd (120˚) | 3rd (240˚) | Mean Value
| 0 | 0 | 0 |  |  |  | 0 | 0 | 0
| 0 | 0 | 0
Rel std unc of force measuring instrument wstd
8.) Uncertainty Calculations
Relative Expanded Uncertainty | Calibration Points | Relative uncertainty by rep (wrep,i) | Rel unc due to res (wres,i) | Relative std uncertainty (wstd) | Relative combined unc (wci) | Relative expanded unc (Wi) | Relative accu error qi | Relative rep error bi | Relative zero error f0 | Class | Calibration uncertainty (wstdc) | Uncertainty due to temp dif (wtemp') | Uncertainty due to sen drift (wdrift) | Relative accuracy error qi | Relative repeatability error bi | Relative zero error f0 | Lower Limit of calibration range | Relative uncertainty of force measuring instrument | Class
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0
Resolution | 1st
2nd
temp sens. | 3rd | [12 empty] | - | - | - | - | - | -
0.003 | 4th
Temp Diff. | 5th
6th
Drift Sens. | 7th
0.05 | 8th
9th
10th
Calibrated by: | [5 empty] | Checked by:
RFM |  |  | ACCG
date:

## Certificate 10PTS (hidden)
CALIBRATION CERTIFICATE
Date of Calibration |  |  | :
Calibration Item |  |  | :
| Capacity |  | :
| Measurement Range |  | :
:
| Make / Model |  | :
| Serial No. |  | :
Customer |  |  | :
MEASUREMENT RESULTS:
Refer to the "Calibration Results" on the next page(s).
UNCERTAINTY OF MEASUREMENT:
The uncertainty stated is the expanded uncertainty obtained by multiplying the standard uncertainty by the coverage factor k = 2. It has been determined in accordance with the “JCGM 100:2008 Evaluation of measurement data- Guide to the Expression of Uncertainty in measurement ”. The value of the measurand lies within the assigned range of values with a probability of approximately 95%.
STANDARD USED :
Name of Standard |  |  | Make/Model |  | Calibration Certificate No. |  |  | Traceability
Traceable to SI through KRISS - Korea
CALIBRATION PROCEDURE:
The testing machine under calibration was subjected to specified force values which were applied on the force-proving instrument and readings were noted.
The relevant method for this calibration is ISO 7500-1:2018, Metallic materials – Verification of static uniaxial testing machines.
ENVIRONMENTAL CONDITIONS:
Ambient Temperature
Relative Humidity
Page 1 of 2
CALIBRATION RESULTS
Compression
Indicated Force |  |  | Applied Force | Relative Expanded Uncertainty |  | Relative Accuracy Error | Relative Repeatability Error |  | Class
|  | kN | kN | % |  | % | %
|  |  |  | 0 |  | 0 | 0
REMARKS:
- | The above results were those obtained at the time of calibration and refer only to the testing machine calibrated. | [8 empty] | The testing machine was found to comply with the requirements to the maximum permissible values (%) for the different relative errors of the force measuring system which characterizes a testing machine in accordance with the appropriate class as given in Table 2, Item 6 of the above method.
-
- | No adjustment was performed on the testing machine. The user should determine suitability of the testing machine for its intended use.
-
-
AHDRIAN CAMILO C. GERNALE
Science Research Specialist I
RADLEY F. MANALO
Science Research Specialist II
For the Chief, National Metrology Laboratory
MARYNESS I. SALAZAR
Section Head
Page 2 of 2

## Certificate 10 pts (hidden)
CALIBRATION CERTIFICATE
Date of Calibration |  |  | :
Calibration Item |  |  | :
| Capacity |  | :
| Measurement Range |  | :
:
| Make / Model |  | :
| Serial No. |  | :
Customer |  |  | :
MEASUREMENT RESULTS:
Indicated Force |  |  | Applied Force | Relative Expanded Uncertainty |  | Relative Accuracy Error | Relative Repeatability Error |  | Class
|  | kN | kN | % |  | % | %
|  |  |  | 0 |  | 0 | 0
|  |  |  | 0 |  | 0 | 0
UNCERTAINTY OF MEASUREMENT:
The uncertainty stated is the expanded uncertainty obtained by multiplying the standard uncertainty by the coverage factor k = 2. It has been determined in accordance with the “JCGM 100:2008 Evaluation of measurement data- Guide to the Expression of Uncertainty in measurement ”. The value of the measurand lies within the assigned range of values with a probability of approximately 95%.
Page 1 of 2
STANDARD USED :
Name of Standard |  |  | Make/Model |  | Calibration Certificate No. |  |  | Traceability
Traceable to SI through KRISS - Korea
CALIBRATION PROCEDURE:
The testing machine under calibration was subjected to specified force values which were applied on the force-proving instrument and readings were noted.
The relevant reference document for this calibration is ISO 7500-1:2018, Metallic materials – Verification of static uniaxial testing machines.
ENVIRONMENTAL CONDITIONS:
Ambient Temperature
Relative Humidity
REMARKS:
- | [9 empty] | The testing machine was found to comply with the requirements to the maximum permissible values (%) for the different relative errors of the force measuring system which characterizes a testing machine in accordance with the appropriate class as given in Table 2, Item 6 of the above method.
-
- | No adjustment was performed on the testing machine. The user should determine suitability of the testing machine for its intended use.
-
-
-
-
Page 2 of 2

## Certificate 5 pts (SI)
CALIBRATION CERTIFICATE
Date of Calibration |  |  | :
Calibration Item |  |  | :
| Capacity |  | :
| Measurement Range |  | :
:
| Make / Model |  | :
| Serial No. |  | :
Customer |  |  | :
MEASUREMENT RESULTS:
|  | Applied Force (F std ) | Indicated Force (F ind) |  | Deviation (F ind - F std) |  | Relative Expanded Uncertainty |  | Relative Accuracy Error
|  | kgf | kgf |  | kgf |  | % |  | %
|  | 0 | 0 |  | 0
UNCERTAINTY OF MEASUREMENT:
The uncertainty stated is the expanded uncertainty obtained by multiplying the standard uncertainty by the coverage factor k = 2. It has been determined in accordance with the “JCGM 100:2008 Evaluation of measurement data- Guide to the expression of uncertainty in measurement”. The value of the measurand lies within the assigned range of values with a probability of approximately 95%.
Page 1 of 2
STANDARD USED :
Name of Standard |  |  |  | Make/Model |  | Calibration Certificate No. |  |  |  | Traceability
[6 empty] | 02-2025-FORC-0072 |  |  |  | Traceable to the SI through NMD-ITDI
CALIBRATION PROCEDURE:
The axle weighing scale was subjected to specified force values in comparison with force reference standard values. Three (3) series of increasing force values were applied to the axle weighing scale.
The relevant references for this axle weighing scale calibration are the ISO 7500-1:2018 "Metallic materials-Calibration and verification of static uniaxial testing machine'' and TP-PFSS-F-FORC-02: (version 1) "Calibration of Axle Weighing Scales" .
ENVIRONMENTAL CONDITIONS:
Ambient Temperature
Relative Humidity
REMARKS:
-
-
-
-
2026-05-19
-End of Report-
Page 2 of 2

## Records
Form No. F5.5.2
Version 2 Page 1 of 1
CALIBRATION RECORD
(Use as first page of Data Sheet for any calibration or measurement work performed.)
Request Reference No.:
Instrument:
Make / Model No.
Customer:
Environmental Conditions:
Calibration Procedure: |  |  | ISO 7500/1
Date of Calibration: |  | 2018-02-21 |  |  | Calibrated by: |  | RFM
Standards and Equipment Used:
Description |  | Make / Model No. |  | Serial No. |  |  |  | Date of Next Calibration
working standard (UTM)
Thermohygrometer |  | TandD/TR-73U |  | 104-104
Remarks:
Checked by:

## Others db
Types of Machine | Name | Position | Unit | equivalence to kN
Box Compression Machine | Ahdrian Camilo C. Gernale | Science Research Specialist II | kgf | 0.0098067
Axle Weighing Scale | Radley F. Manalo | Senior Science Research Specialist | lbf | 0.0044482
Universal Testing Machine | Glenn Joseph C. Franes | Science Research Specialist I | tonf
Edge Crush Testing Machine | Maryness I. Salazar | Section Head | n | 0.0001 | 1000 | kg | 1 | ton
Ring Crush Testing Machine | Pierre Jordan M. Mendoza | Science Research Specialist I
Compression Testing Machine
Tensile Testing Machine
Force Measuring Instrument

## 123 (hidden)
Force, N
0 | 0
196.133 | 0.19612
392.266 | 0.39222
588.399 | 0.58832
748.532 | 0.78444
980.665 | 0.98056
1176.798 | 1.1767
1569.064 | 1.3728
1765.197 | 1.569
1961.33 | 1.7651