## Dash Board
[17 empty] | Equipment Photo: |  | Photo 1
Name/Company : | SAVVICE CORPORATION |  |  |  |  | Place of Calibration : |  | Force Standards Laboratory, NML ITDI |  |  |  |  | Machine Indication
Address Line 1 : | 3rd Floor BMWC Bldg., Cagayan Valley Road | [11 empty] | 0 | 0
Address Line 2 : | Sta. Rita, Guiginto, Bulacan |  |  |  |  | Person in Charge : | [6 empty] | 1st | 3000
Requesting Official : | Ramil F. De Luna |  |  |  |  | Analyst/Calibrator : |  | Ahdrian Camilo C. Gernale |  |  |  |  | 2nd | 6000
Position : | Technician | [6 empty] | Radley F. Manalo |  |  |  |  | 3rd | 9000
[13 empty] | 4th | 12000
[13 empty] | 5th | 15000
6th
[13 empty] | 7th | [5 empty] | Photo 2
8th
Date of Calibration : | 2026-04-22 |  |  |  |  | Calibration Mode |  | Compression |  |  |  |  | 9th
Req. Reference No. : | 03-2026-FORC-0087 |  |  |  |  | Capacity : |  | 15000 | kgf |  |  |  | 10th
Calibration Item : | Axle Weighing Scale |  |  |  |  | Range : |  | 0 | kgf to | 15000 | kgf |  | 0 | 0
Calibration Type : | [5 empty] | Increment | : | 50 | kgf
Load Cell Type : | No |  |  |  |  | Resolution : |  | 50 | kgf
Force Transducer Info
Make / Model No. : | LS 630 W. PAD
Serial No. / ID No. : | 0122YL21002
Indicator Info
Make / Model No. : | INTERCOMP PT20
Serial No. / ID No. : | 0307YM23003
Special Instructions & Remarks :
please input above the preparations and other pertinent details for the calibrators to remember before starting
Calibrator Tools :
TSR # | Date of Calibration | In-Charge/ Supervisor |  | Location | Standards Used | Calibrator/ Analyst | Test Points | [9 empty] | Equivalent Force | [9 empty] | U | Class | Remarks
[7 empty] | 1st | 2nd | 3rd | 4th | 5th | 6th | 7th | 8th | 9th | 10th | 1st | 2nd | 3rd | 4th | 5th | 6th | 7th | 8th | 9th | 10th
Note on D7: This will be displayed in the certificate
Force Standards Section
B. EQUIPMENT INFORMATION
C. CALIBRATION HISTORY
A. CUSTOMER'S INFORMATION
D. CALIBRATION POINTS
New Calibration
Print Certificate
Print Cal Data
Print Cal Record

## Software (hidden)
FORCE MEASUREMENT DATA
Ver. 2 Rev.0 Mar 12, 2019
revised & developed by: gearnhale@yahoo.com according to ISO
original developer: manuelmruiz@yahoo.com
Calibration of Standard Expired - DO NOT USE!
Date of Calibration | 2026-04-22 | Calibration Mode |  |  | Compression
Request Reference No. | 03-2026-FORC-0087 | Capacity |  |  | 15000
Calibration Item | Axle Weighing Scale | Range |  |  | 0 kgf to 15000 kgf
Make / Model No. | LS 630 W. PAD | Increment/Graduation |  |  | 50 kgf
Serial No. / ID No. | 0122YL21002 | Resolution |  |  | 50 kgf
Guide on how to use this data entry system: |  |  | Environmental Conditions:
1. Select the radio button of the machine indicator unit (kgf, lbf, N, kN). |  |  |  | Before | After
2. After capturing the digital amplifier's reading, change the machine indication accordingly. |  |  | Temp: | 25 | 25
3. Upon completion of each loading set, press the corresponding PRINT button. Make sure you have |  |  | Humidity: | 33 | 33
change first the machine indication reading upon capturing.
Data Logger |  |  | Calibration Worksheet
[8 empty] | Machine Indicator Readability |  | 50 kgf
Zero | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN |  | : Pre-load from 0 to 100
| Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0 |  | Point No. | kgf | kN |  |  | raw | zero corrected
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0
1st Test point | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | 2nd | 40 | 0.39227 | 1.1 1.1-DMP40 -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.24268 | -0.32077 | -1.5727
| Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0 |  | 3rd | 60 | 0.5884 | 1.1 1.1-DMP40 -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.40344 | -0.48153 | -2.361
|  |  | 4th | 80 | 0.78453 | 1.1 1.1-DMP40 -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.56347 | -0.64155 | -3.1457
2nd Test point | 1.1 1.1-DMP40 -0.242682 mV/V abs 13-Mar-19 13:06:55 |  | 5th | 100 | 0.98067 | 1.1 1.1-DMP40 -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.72362 | -0.80171 | -3.9311 |  | : Increase load 0 to 0
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | 6th |  | - - | 0 | #VALUE! | - - | - - | - -
| Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0 |  | 7th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 8th |  | - - | 0 | #VALUE! | - - | - - | - -
3rd Test point | 1.1 1.1-DMP40 -0.403444 mV/V abs 13-Mar-19 13:07:17 |  | 9th |  | - - | 0 | #VALUE! | - - | - - | - -
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | 10th |  | - - | 0 | #VALUE! | - - | - - | - -
| Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0 |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475
Force Transducer Information
4th Test point | 1.1 1.1-DMP40 -0.563466 mV/V abs 13-Mar-19 13:08:12 | [8 empty] | 0.0098067 | kgf
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | Load Cell | HBM/C3H3 |  |  | 6 | A | 4.9026
| Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0 |  | Capacity | 10 kN |  |  | 7 | B | -0.00093075
|  |  | Serial No. | F93007 |  |  | 8 | C | -6.3551e-05
5th Test point | 1.1 1.1-DMP40 -0.723624 mV/V abs 13-Mar-19 13:09:02 |  | Cal. Cert. No. | 1800-00040-004 |  |  | 9 | U, % | 0.011
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | Cal Date | 2018-03-02 |  | 2021-03-02 | 10 | k | 2
Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0
6th Test point |  |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
|  |  | Point No. | kgf | kN |  |  | raw | zero corrected
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0
|  |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971
7th Test point |  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475
Pre-Load 2
|  |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
8th Test point |  |  | Point No. | kgf | kN |  |  | raw | zero corrected
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0
|  |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475
9th Test point | [13 empty] | 3
Pre-Load 3
|  |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
|  |  | Point No. | kgf | kN |  |  | raw | zero corrected
10th Test point |  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0
|  |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475
Zero | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 |  | Series 1
| Excitation: 5.0 V Meas. Range: 2.5 mV/V |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
| Meas. Filt: 0.90 Hz Be AutoCal: Off LV:1=0,2=0,3=0,4=0 |  | Point No. | kgf | kN |  |  | raw | zero corrected
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0 |  | Error Estimation, in % | 1000
|  |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971 |  | -124.8361
|  |  | 2nd | 40 | 0.39227 | 1.1 1.1-DMP40 -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.24268 | -0.32077 | -1.5727 |  | -124.9422
|  |  | 3rd | 60 | 0.5884 | 1.1 1.1-DMP40 -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.40344 | -0.48153 | -2.361 |  | -124.9219
|  |  | 4th | 80 | 0.78453 | 1.1 1.1-DMP40 -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.56347 | -0.64155 | -3.1457 |  | -124.9401
|  |  | 5th | 100 | 0.98067 | 1.1 1.1-DMP40 -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.72362 | -0.80171 | -3.9311 |  | -124.9466
|  |  | 6th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 7th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 8th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 9th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 10th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475
Series 2
|  |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
|  |  | Point No. | kgf | kN |  |  | raw | zero corrected
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0 |  | Error Estimation, in %
|  |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971 |  | -124.8361
|  |  | 2nd | 40 | 0.39227 | 1.1 1.1-DMP40 -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.24268 | -0.32077 | -1.5727 |  | -124.9422
|  |  | 3rd | 60 | 0.5884 | 1.1 1.1-DMP40 -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.40344 | -0.48153 | -2.361 |  | -124.9219
|  |  | 4th | 80 | 0.78453 | 1.1 1.1-DMP40 -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.56347 | -0.64155 | -3.1457 |  | -124.9401
|  |  | 5th | 100 | 0.98067 | 1.1 1.1-DMP40 -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.72362 | -0.80171 | -3.9311 |  | -124.9466
|  |  | 6th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 7th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 8th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 9th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 10th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475
Series 3
|  |  | Test | Machine Indication |  |  |  | Load Cell Indication, mV/V |  | Equivalent Force, kN
|  |  | Point No. | kgf | kN |  |  | raw | zero corrected
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 mV/V abs 13-Mar-19 13:04:56 | 0.078086 | 0 | 0 |  | Error Estimation, in %
|  |  | 1st | 20 | 0.19613 | 1.1 1.1-DMP40 -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 mV/V abs 13-Mar-19 13:06:28 | -0.082988 | -0.16107 | -0.78971 |  | -124.8361
|  |  | 2nd | 40 | 0.39227 | 1.1 1.1-DMP40 -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.242682 mV/V abs 13-Mar-19 13:06:55 | -0.24268 | -0.32077 | -1.5727 |  | -124.9422
|  |  | 3rd | 60 | 0.5884 | 1.1 1.1-DMP40 -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.403444 mV/V abs 13-Mar-19 13:07:17 | -0.40344 | -0.48153 | -2.361 |  | -124.9219
|  |  | 4th | 80 | 0.78453 | 1.1 1.1-DMP40 -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.563466 mV/V abs 13-Mar-19 13:08:12 | -0.56347 | -0.64155 | -3.1457 |  | -124.9401
|  |  | 5th | 100 | 0.98067 | 1.1 1.1-DMP40 -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.723624 mV/V abs 13-Mar-19 13:09:02 | -0.72362 | -0.80171 | -3.9311 |  | -124.9466
|  |  | 6th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 7th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 8th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 9th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | 10th |  | - - | 0 | #VALUE! | - - | - - | - -
|  |  | Zero | 0 | 0 | 1.1 1.1-DMP40 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 mV/V abs 13-Mar-19 13:10:36 | 0.077444 | -0.000642 | -0.0031475
STEP 1
STEP 2
PRINT Series 1
PRINT Series 2
PRINT Series 3
ERASE Pre-Load 3
ERASE Series 1
ERASE Series 2
kgf
lbf
N
kN
PRINT Pre-Load 1
ERASE Series 3
PRINT Pre-Load 2
PROCESS RESULT CONTROLLER
DELETE Controller
PRINT Pre-Load 3
ERASE Pre-Load 1
ERASE Pre-Load 2
ERASE ALL LOAD

## Calibration Data (hidden)
FORCE MEASUREMENT DATA
Ver. 2 Rev.0 Mar 12, 2019
revised & developed by: gearnhale@yahoo.com according to ISO
original developer: manuelmruiz@yahoo.com
Certificate No. |  | 03-2026-FORC-0087
Client Name: |  | SAVVICE CORPORATION
Calibration Date: |  | 2026-04-22
Client Address: |  | 3rd Floor BMWC Bldg., Cagayan Valley Road Sta. Rita, Guiginto, Bulacan | [6 empty] | Calibration Mode: |  | Compression
Calibration Item: |  | Axle Weighing Scale | [6 empty] | Machine Capacity: |  | 15000 kgf
Model Code/No: |  | LS 630 W. PAD | [6 empty] | Calibration Range: |  | 0 kgf to 15000 kgf
Serial No/ ID No. |  | 0122YL21002 | [6 empty] | Increment |  | 50
Remarks: | [8 empty] | Resolution
1. Preloading
|  | Indicated Force | Load 1st (0˚) | Indicated Force | Load 2nd (120˚) | Indicated Force | Load 3rd (240˚) |  | Reference Standard Used: |  | HBM/C3H3 |  |  | Temp.: | Humidity:
|  | 0 | 0.078086 | 0 | 0.078086 | 0 | 0.078086 |  | Capacity: 10 kN |  | a' | 4.9026 |  | 25 | 33
|  | 0.19613 | -0.082988 | 0.19613 | -0.082988 | 0.19613 | -0.082988 |  | Serial #: F93007 |  | b' | -0.00093075 |  | 25 | 33 | [10 empty] | Relative Expanded Uncertainty |  |  | Relative standard uncertainty of force measuring instrument wstd | [13 empty] | #DIV/0!
|  | 0 | 0.077444 | 0 | 0.077444 | 0 | 0.077444 |  | Cal Cert #: 1800-00040-004 |  | c' | -6.3551e-05 |  | 25 | 33
[18 empty] | RAW DATA CALCULATION | [8 empty] | Relative uncertainty by reproducibility (wrep,i) | Relative uncertainty due to resolution (wres,i) | Calibration uncertainty (wstdc) | Uncertainty due to temperature difference (wtemp') | Uncertainty due to sensitivity drift (wdrift) | Relative standard uncertainty (wstd) | Relative combined uncertainty (wci) | Effective degree of freedom | coverage factor (k) | Relative expanded uncertainty (Wi) | Relative accu error qi | Relative rep error bi | Relative zero error f0 | Relative accuracy error qi | Relative repeatability error bi | Relative zero error f0 | Lower Limit of calibration range | Relative uncertainty of force measuring instrument 상대불확도 | Class
| 2. Linear Interpolation (of reference force deflection) | [6 empty] | 3. Estimation of reference force | [9 empty] | Net Value (dij) & Mean (di) Calculation | [7 empty] | 0.011
|  | Force (kN) | Load Value | [5 empty] | Force (kN) | Load Value (kN) | [5 empty] | Machine Indication |  | Indicated Force | Load 1st (0˚) | Indicated Force | Load 2nd (120˚) | Indicated Force | Load 3rd (240˚) | Mean Value | Applied Force | Resolution testing machine
|  |  | 1st (0˚) | 2nd (120˚) | 3rd (240˚) | Mean Value |  |  |  | 1st (0˚) | 2nd (120˚) | 3rd (240˚) | Mean Value |  |  | Indicated Force | Equivalent Force (kN)
|  | 0 | 0 | 0 | 0 | 0 |  |  | 0 | -0.004913 | 0 | -0.004913 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | - | - | 0 | 0 | 0 | 0 | - | - | - | - | - | -
|  | 29.42 | -24.1611 | -24.1611 | -24.1611 | -24.1611 |  |  | 29.42 | 118.1 | 118.1 | 118.1 | 118.1 |  |  | 3000 | 29.42 | 0.19613 | -0.16107 | 0.19613 | -0.16107 | 0.19613 | -0.16107 | -0.16107 | 118.1 | Sensitivity on temp | 0 | 0 | 0.0055 | 0 | 0.028868 | 0.029387 | 0.029387 | #DIV/0! | 2 | 0.058774 | -75.089 | 0 | 0 | - | 0.5 | 0.5 | #DIV/0! | 0.5 | 0.5
|  | 58.8399 | -48.1152 | -48.1152 | -48.1152 | -48.1152 |  |  | 58.8399 | 230.967 | 230.967 | 230.967 | 230.967 |  |  | 6000 | 58.8399 | 0.39227 | -0.32077 | 0.39227 | -0.32077 | 0.39227 | -0.32077 | -0.32077 | 230.967 |  | 0 | 0 | 0.0055 | 0 | 0.028868 | 0.029387 | 0.029387 | #DIV/0! | 2 | 0.058774 | -74.5245 | 0 | 0 | - | 0.5 | 0.5 | #DIV/0! | 0.5 | 0.5
|  | 88.2599 | -72.2295 | -72.2295 | -72.2295 | -72.2295 |  |  | 88.2599 | 335.023 | 335.023 | 335.023 | 335.023 |  |  | 9000 | 88.2599 | 0.5884 | -0.48153 | 0.5884 | -0.48153 | 0.5884 | -0.48153 | -0.48153 | 335.023 | 0.003 | 0 | 0 | 0.0055 | 0 | 0.028868 | 0.029387 | 0.029387 | #DIV/0! | 2 | 0.058774 | -73.6556 | 0 | 0 | - | 0.5 | 0.5 | #DIV/0! | 0.5 | 0.5
|  | 117.6798 | -96.2328 | -96.2328 | -96.2328 | -96.2328 |  |  | 117.6798 | 423.778 | 423.778 | 423.778 | 423.778 |  |  | 12000 | 117.6798 | 0.78453 | -0.64155 | 0.78453 | -0.64155 | 0.78453 | -0.64155 | -0.64155 | 423.778 | Temp Diff. | 0 | 0 | 0.0055 | 0 | 0.028868 | 0.029387 | 0.029387 | #DIV/0! | 2 | 0.058774 | -72.2308 | 0 | 0 | - | 0.5 | 0.5 | #DIV/0! | 0.5 | 0.5
|  | 147.0998 | -120.2565 | -120.2565 | -120.2565 | -120.2565 |  |  | 147.0998 | 492.512 | 492.512 | 492.512 | 492.512 |  |  | 15000 | 147.0998 | 0.98067 | -0.80171 | 0.98067 | -0.80171 | 0.98067 | -0.80171 | -0.80171 | 492.512 | 0 | 0 | 0 | 0.0055 | 0 | 0.028868 | 0.029387 | 0.029387 | #DIV/0! | 2 | 0.058774 | -70.1328 | 0 | 0 | - | 0.5 | 0.5 | #DIV/0! | 0.5 | 0.5
|  | 0 | - - | - - | - - | - - |  |  | 0 | - - | - - | - - | - - |  |  | 0 | 0 | - - | - - | - - | - - | - - | - - | - - | - - | Drift Sens. | #VALUE! | #VALUE! | 0.0055 | 0 | 0.028868 | 0.029387 | #VALUE! |  | 2 | #VALUE! | - - | - - | 0 | #VALUE! | - | 0.5 | #DIV/0! | 0.5 | #VALUE!
|  | 0 | - - | - - | - - | - - |  |  | 0 | - - | - - | - - | - - |  |  | 0 | 0 | - - | - - | - - | - - | - - | - - | - - | - - | 0.05 | #VALUE! | #VALUE! | 0.0055 | 0 | 0.028868 | 0.029387 | #VALUE! |  | 2 | #VALUE! | - - | - - | 0 | #VALUE! | - | 0.5 | #DIV/0! | 0.5 | #VALUE!
|  | 0 | - - | - - | - - | - - |  |  | 0 | - - | - - | - - | - - |  |  | 0 | 0 | - - | - - | - - | - - | - - | - - | - - | - - |  | #VALUE! | #VALUE! | 0.0055 | 0 | 0.028868 | 0.029387 | #VALUE! |  | 2 | #VALUE! | - - | - - | 0 | #VALUE! | - | 0.5 | #DIV/0! | 0.5 | #VALUE!
|  | 0 | - - | - - | - - | - - |  |  | 0 | - - | - - | - - | - - |  |  | 0 | 0 | - - | - - | - - | - - | - - | - - | - - | - - |  | #VALUE! | #VALUE! | 0.0055 | 0 | 0.028868 | 0.029387 | #VALUE! |  | 2 | #VALUE! | - - | - - | 0 | #VALUE! | - | 0.5 | #DIV/0! | 0.5 | #VALUE!
|  | 0 | - - | - - | - - | - - |  |  | 0 | - - | - - | - - | - - |  |  | 0 | 0 | - - | - - | - - | - - | - - | - - | - - | - - |  | #VALUE! | #VALUE! | 0.0055 | 0 | 0.028868 | 0.029387 | #VALUE! |  | 2 | #VALUE! | - - | - - | 0 | #VALUE! | - | 0.5 | #DIV/0! | 0.5 | #VALUE!
|  | 0 | 0 | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 |  |  | 0 | 0 | 0 | -0.000642 | 0 | -0.000642 | 0 | -0.000642 | -0.00064 | 0
[16 empty] | Max Zero | = | 0
1.0049e-14
Pressure Force Standards Section
National Metrology Laboratory of the Philippines

## Load Cell db (hidden)
Description | Capacity | Serial Number | Calibration Certificate | Calibration Date | a' (Compression) | b' (Compression) | c' (Compression) | U (Compression) | a' (Tension) | b' (Tension) | c' (Tension) | U (Tension) | Next Cal Date
HBM/C6 | 2000 kN | F42442 | 1800-00040-002 | 2018-02-22 | 978.3389 | 35.0482 | -7.9402 | 1.375 | -- | -- | -- | -- | 2021-02-22
HBM/TCBS | 500 kN | 50036 | 1800-00040-003 | 2018-03-02 | -247.8101 | 0.47744 | -0.010141 | 0.07 | -- | -- | -- | -- | 2021-03-02
HBM/C3H3 | 10 kN | F93007 | 1800-00040-004 | 2018-03-02 | 4.9026 | -0.00093075 | -6.3551e-05 | 0.011 | -4.9023 | 0.00041063 | -0.00016826 | 0.012 | 2021-03-02
HBM/Z3H3 | 5 kN | F73724 | 1800-00040-005 | 2018-03-02 | 2.4495 | 0.00032711 | -0.00017387 | 0.017 | -2.4504 | -0.00031491 | -9.9356e-05 | 0.019 | 2021-03-02
HBM/Z3OA | 2 kN | 151730166 | 1600-02019-001 | 2016-06-17 | 1000.037 | -0.091009 | -0.0017968 | 0.024 | -- | -- | -- | -- | 2019-06-17
HBM/Z4A (20kN) | 20 kN | 161430192 | 1600-02019-003 | 2016-06-24 | 10.0026 | 0.00072857 | -0.00058347 | 0.019 | -10.0022 | 0.0018933 | 0.00041544 | 0.012 | 2019-06-24
HBM/Z4A (100kN) | 100 kN | 160730020 | 1600-02019-002 | 2016-07-29 | 50.0339 | -0.014762 | -0.0022079 | 0.041 | -50.0346 | 0.0062677 | 0.0026657 | 0.053 | 2019-07-29
WBRC/R200K | 200 kN | WB-1210-010 | 1600-02019-004 | 2016-07-25 | 90.7829 | -0.032822 | -0.0048756 | 0.081 | -- | -- | -- | -- | 2019-07-25
WBRC-R2M | 2 MN | WB-121005 | 1800-00040-001 | 2018-02-22 | 1379.299 | 24.7927 | 18.2805 | 0.078 | -- | -- | -- | -- | 2021-02-22
Shimadzu/ UH-F1000kNX | 1000 kN | 1251056K0094 | 11-2020-FORC-0116 | 2020-11-03 |  |  |  | 0.23 | -- | -- | -- | 0.1 | 2023-11-03
Formulas:
Next Cal Date = DATE(YEAR(Table1[[#This Row],[Calibration Date]])+3, MONTH(Table1[[#This Row],[Calibration Date]]),DAY(Table1[[#This Row],[Calibration Date]]))

## Data
CALIBRATION DATA SHEET
Customer Name: |  | - -
Address Line 1: |  | - -
Address Line 2: |  | - -
Date of Calibration: |  | - - | [5 empty] | Mode: | Compression
Request Ref. No.: |  | Dynamic Unit Test Conv | [5 empty] | Capacity: | 100 tf
Calibration Item: |  | - - | [5 empty] | Range: | - -
Load Cell
Make / Model No.: |  | - - | [5 empty] | Increment: | - -
Serial No. / ID No.: |  | - - | [5 empty] | Resolution: | - -
Indicator
Make / Model No.: |  | - -
Serial No. / ID No.: |  | - -
2.) PRE-LOADING DATA, tf
Indicated force | 1st (0˚) | Indicated force |  | 2nd (120˚) | Indicated force | 3rd (240˚) |  | Conversion Unit:
0 | 0 | 0 |  |  | 0 |  |  | 1 tf = | 9.8066 | kN
15000 | 14950 | 15000 |  |  | 15000
0 | 0 | 0 |  |  | 0
3.) MEASURED DATA, tf
| Target (tf) | [7 empty] | Mean Force (tf)
Calibration Points | Expected Machine Indication | Indicated force |  | 1st | Indicated force | 2nd | Indicated force | 3rd | Mean Force | Mean deflection
0 | - - | - - |  | - - | - - | - - | - - | - - | - - | - -
1st | - - | - - |  | - - | - - | - - | - - | - - | - - | - -
2nd | - - | - - |  | - - | - - | - - | - - | - - | - - | - -
3rd | - - | - - |  | - - | - - | - - | - - | - - | - - | - -
4th | - - | - - |  | - - | - - | - - | - - | - - | - - | - -
5th | - - | - - |  | - - | - - | - - | - - | - - | - - | - -
6th | - - | - - |  | - - | - - | - - | - - | - - | - - | - -
7th | - - | - - |  | - - | - - | - - | - - | - - | - - | - -
8th | - - | - - |  | - - | - - | - - | - - | - - | - - | - -
9th | - - | - - |  | - - | - - | - - | - - | - - | - - | - -
10th | - - | - - |  | - - | - - | - - | - - | - - | - - | - -
0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
Max Zero =
3.) Force Transducer Information | [7 empty] | 4.) Environmental Conditions:
In: Compression Mode | [7 empty] | Thermo-hygrometer indication
Load Cell | - - |  | 6 | A | - - | equalsVLOOKUP($E$36,Table1,H36,FALSE) |  |  | Before | After
Capacity | #N/A |  | 7 | B | - - | equalsVLOOKUP($E$36,Table1,H37,FALSE) |  | Temp: | - - | - -
Serial No. | - - |  | 8 | C | - - | equalsVLOOKUP($E$36,Table1,H38,FALSE) |  | Humidity: | - - | - -
Cal. Cert. No. | - - |  | 9 | U, % | - - | equalsVLOOKUP($E$36,Table1,H39,FALSE)
Cal Date | - - |  | 1900-01-09 | k | 2
Calibrated by: | [6 empty] | Checked by:
RFM |  |  |  | ACCG
date:
5.) Net value (dij) & Mean Value (di) Calculation, units in tf
Calibration Points | Force (tf) | 1st (0˚) |  | Force (tf) | 2nd (120˚) | Force (tf) | 3rd (240˚) | Mean Force (tf) | Mean deflection
0 | 0 | - - |  | 0 | - - | 0 | - - | - - | - -
1st | 0 | - - |  | 0 | - - | 0 | - - | - - | - -
2nd | 0 | - - |  | 0 | - - | 0 | - - | - - | - -
3rd | 0 | - - |  | 0 | - - | 0 | - - | - - | - -
4th | 0 | - - |  | 0 | - - | 0 | - - | - - | - -
5th | 0 | - - |  | 0 | - - | 0 | - - | - - | - -
6th | 0 | - - |  | 0 | - - | 0 | - - | - - | - -
7th | 0 | - - |  | 0 | - - | 0 | - - | - - | - -
8th | 0 | - - |  | 0 | - - | 0 | - - | - - | - -
9th | 0 | - - |  | 0 | - - | 0 | - - | - - | - -
10th | 0 | - - |  | 0 | - - | 0 | - - | - - | - -
0 | 0 | #VALUE! |  | 0 | #VALUE! | 0 | #VALUE!
6.) Linear Interpolation | [6 empty] | 7. Estimation of Reference Force
Force (tf) | Load Value | [5 empty] | Force (tf) | Load Value
| 1st (0˚) | 2nd (120˚) |  | 3rd (240˚) | Mean Value |  |  | Force 1 (tf) | Force 2 (tf) | Force 3 (tf) | Mean Force (tf)
- - | - - | - - |  | - - | - - |  | - - | - - | - - | - - | - -
- - | - - | - - |  | - - | - - |  | - - | - - | - - | - - | - -
- - | - - | - - |  | - - | - - |  | - - | - - | - - | - - | - -
- - | - - | - - |  | - - | - - |  | - - | - - | - - | - - | - -
- - | - - | - - |  | - - | - - |  | - - | - - | - - | - - | - -
- - | - - | - - |  | - - | - - |  | - - | - - | - - | - - | - -
- - | - - | - - |  | - - | - - |  | - - | - - | - - | - - | - -
- - | - - | - - |  | - - | - - |  | - - | - - | - - | - - | - -
- - | - - | - - |  | - - | - - |  | - - | - - | - - | - - | - -
- - | - - | - - |  | - - | - - |  | - - | - - | - - | - - | - -
- - | - - | - - |  | - - | - - |  | - - | - - | - - | - - | - -
0 | 0 | 0 |  | 0 | 0 |  | 0 | - - | - - | - - | - -
[12 empty] | Rel std unc of force measuring instrument wstd | [5 empty] | #VALUE!
8.) Uncertainty Calculations
Relative Expanded Uncertainty | Calibration Points | Relative uncertainty by rep (wrep,i) |  | Rel unc due to res (wres,i) | Relative std uncertainty (wstd) | Relative combined unc (wci) | Relative expanded unc (Wi) | Relative accu error qi | Relative rep error bi | Relative zero error f0 | Class | Calibration uncertainty (wstdc) | Uncertainty due to temp dif (wtemp') | Uncertainty due to sen drift (wdrift) | Relative accuracy error qi | Relative repeatability error bi | Relative zero error f0 | Lower Limit of calibration range | Relative uncertainty of force measuring instrument | Class
- - | 0 | - - |  | - - | - - | - - | - - | - - | - - | - - | - - | 0 | 0 | 0
Resolution | 1st | - - |  | - - | - - | - - | - - | - - | - - | - - | - - | #VALUE! | #VALUE! | 0.028868
490.3325 | 2nd | - - |  | - - | - - | - - | - - | - - | - - | - - | - - | #VALUE! | #VALUE! | 0.028868
temp sens. | 3rd | - - |  | - - | - - | - - | - - | - - | - - | - - | - - | #VALUE! | #VALUE! | 0.028868 | - | - | - | - | - | -
0.003 | 4th | - - |  | - - | - - | - - | - - | - - | - - | - - | - - | #VALUE! | #VALUE! | 0.028868 | - - | - - | - - | - - | - - | - -
Temp Diff. | 5th | - - |  | - - | - - | - - | - - | - - | - - | - - | - - | #VALUE! | #VALUE! | 0.028868 | - - | - - | - - | - - | - - | - -
#VALUE! | 6th | - - |  | - - | - - | - - | - - | - - | - - | - - | - - | #VALUE! | #VALUE! | 0.028868 | - - | - - | - - | - - | - - | - -
Drift Sens. | 7th | - - |  | - - | - - | - - | - - | - - | - - | - - | - - | #VALUE! | #VALUE! | 0.028868 | - - | - - | - - | - - | - - | - -
0.05 | 8th | - - |  | - - | - - | - - | - - | - - | - - | - - | - - | #VALUE! | #VALUE! | 0.028868 | - - | - - | - - | - - | - - | - -
| 9th | - - |  | - - | - - | - - | - - | - - | - - | - - | - - | #VALUE! | #VALUE! | 0.028868 | - - | - - | - - | - - | - - | - -
| 10th | - - |  | - - | - - | - - | - - | - - | - - | - - | - - | #VALUE! | #VALUE! | 0.028868 | - - | - - | - - | - - | - - | - -
[11 empty] | 0 |  |  |  | - - | - - | - - | - - | - - | - -
Calibrated by: | [6 empty] | Checked by: | [7 empty] | - - | - - | - - | - - | - - | - -
[15 empty] | - - | - - | - - | - - | - - | - -
RFM |  |  |  | ACCG
date:
no corrections
with corrections
no corrections
with corrections

## Certificate 10PTS (hidden)
CALIBRATION CERTIFICATE
No. 03-2026-FORC-0087
Date of Calibration |  |  | : | 2026-04-22
Calibration Item |  |  | : | Axle Weighing Scale
| Capacity |  | : | 15000 kgf
| Measurement Range |  | : | 0 kgf to 15000 kgf
| Resolution : |  | : | 50
| Make / Model |  | : | LS 630 W. PAD
| Serial No. |  | : | 0122YL21002
(Indicator)
Customer |  |  | : | SAVVICE CORPORATION
3rd Floor BMWC Bldg., Cagayan Valley Road
Sta. Rita, Guiginto, Bulacan
MEASUREMENT RESULTS:
Refer to the "Calibration Results" on the next page(s).
UNCERTAINTY OF MEASUREMENT:
The uncertainty stated is the expanded uncertainty obtained by multiplying the standard uncertainty by the coverage factor k = 2. It has been determined in accordance with the “JCGM 100:2008 Evaluation of measurement data- Guide to the Expression of Uncertainty in measurement ”. The value of the measurand lies within the assigned range of values with a probability of approximately 95%.
STANDARD USED :
Name of Standard |  |  | Make/Model |  |  | Calibration Certificate No. |  |  | Traceability
Load Cell SN F93007 |  |  | HBM/C3H3 |  |  | 1800-00040-004 |  |  | Traceable to SI through KRISS - Korea
CALIBRATION PROCEDURE:
The testing machine under calibration was subjected to specified force values which were applied on the force-proving instrument and readings were noted.
The relevant method for this calibration is ISO 7500-1:2018, Metallic materials – Verification of static uniaxial testing machines.
ENVIRONMENTAL CONDITIONS:
| Ambient Temperature |  |  | : (25± 1) °C
| Relative Humidity |  |  | : (33± 5)
Page 1 of 2
CALIBRATION RESULTS
Compression
Indicated Force |  |  | Applied Force |  | Relative Expanded Uncertainty |  | Relative Accuracy Error | Relative Repeatability Error |  | Class
kgf |  | kN | kN |  | % |  | % | %
0 |  | 0 | 0 |  | 0 |  | 0 | 0 |  | -
3000 |  | 29.42 | 118.1 |  | 0.06 |  | -74.53 | 0 |  | 0.5
6000 |  | 58.8399 | 230.967 |  | 0.06 |  | -73.66 | 0 |  | 0.5
9000 |  | 88.2599 | 335.023 |  | 0.06 |  | -72.24 | 0 |  | 0.5
12000 |  | 117.6798 | 423.778 |  | 0.06 |  | -70.14 | 0 |  | 0.5
15000 |  | 147.0998 | 492.512 |  | #VALUE! |  | #VALUE! | 0 |  | 0.5
0 |  | 0 | - - |  | #VALUE! |  | #VALUE! | #VALUE! |  | #VALUE!
0 |  | 0 | - - |  | #VALUE! |  | #VALUE! | #VALUE! |  | #VALUE!
0 |  | 0 | - - |  | #VALUE! |  | #VALUE! | #VALUE! |  | #VALUE!
0 |  | 0 | - - |  | #VALUE! |  | #VALUE! | #VALUE! |  | #VALUE!
0 |  | 0 | - - |  | 0 |  | 0 | #VALUE! |  | #VALUE!
REMARKS:
- | The above results were those obtained at the time of calibration and refer only to the testing machine calibrated. | [9 empty] | The testing machine was found to comply with the requirements to the maximum permissible values (%) for the different relative errors of the force measuring system which characterizes a testing machine in accordance with the appropriate class as given in Table 2, Item 6 of the above method.
- | The testing machine was found to comply with the requirements to the maximum permissible values (%) for the different relative errors of the force measuring system which characterizes a testing machine in accordance with the appropriate class as given in Table 2, Item 6 of the above method.
- | No adjustment was performed on the testing machine. The user should determine suitability of the testing machine for its intended use.
- | The conversion factor used is: 1 kgf = kN.
- | The testing machine was calibrated at the Force Standards Laboratory, NML ITDI of SAVVICE CORPORATION at 3rd Floor BMWC Bldg., Cagayan Valley Road, Sta. Rita, Guiginto, Bulacan.
AHDRIAN CAMILO C. GERNALE
Science Research Specialist I
RADLEY F. MANALO
Science Research Specialist II
For the Chief, National Metrology Laboratory
MARYNESS I. SALAZAR
Section Head
Page 2 of 2
Calibration Certificate No. 03-2026-FORC-0087

## Certificate 10 pts (hidden)
CALIBRATION CERTIFICATE
No. 03-2026-FORC-0087
Date of Calibration |  |  | : | 2026-04-22
Calibration Item |  |  | : | Axle Weighing Scale
| Capacity |  | : | 15000 kgf
| Measurement Range |  | : | 0 kgf to 15000 kgf
| Resolution : |  | : | 50
| Make / Model |  | : | LS 630 W. PAD (Load Cell); INTERCOMP PT20 (Indicator)
| Serial No. |  | : | 0122YL21002 (Load Cell); 0307YM23003 (Indicator)
(Indicator)
Customer |  |  | : | SAVVICE CORPORATION
3rd Floor BMWC Bldg., Cagayan Valley Road
Sta. Rita, Guiginto, Bulacan
MEASUREMENT RESULTS:
Indicated Force |  |  | Applied Force |  | Relative Expanded Uncertainty |  | Relative Accuracy Error | Relative Repeatability Error |  | Class
kgf |  | kN | kN |  | % |  | % | %
0 |  | 0 | 0 |  | 0 |  | 0 | 0 |  | -
3000 |  | 29.42 | 118.1 |  | 0.06 |  | -75.09 | 0 |  | 0.5
6000 |  | 58.8399 | 230.967 |  | 0.06 |  | -74.53 | 0 |  | 0.5
9000 |  | 88.2599 | 335.023 |  | 0.06 |  | -73.66 | 0 |  | 0.5
12000 |  | 117.6798 | 423.778 |  | 0.06 |  | -72.24 | 0 |  | 0.5
15000 |  | 147.0998 | 492.512 |  | 0.06 |  | -70.14 | 0 |  | 0.5
0 |  | 0 | - - |  | 0 |  | 0 | 0 |  | #VALUE!
0 |  | 0 | - - |  | #VALUE! |  | #VALUE! | #VALUE! |  | #VALUE!
0 |  | 0 | - - |  | #VALUE! |  | #VALUE! | #VALUE! |  | #VALUE!
0 |  | 0 | - - |  | #VALUE! |  | #VALUE! | #VALUE! |  | #VALUE!
0 |  | 0 | - - |  | #VALUE! |  | #VALUE! | #VALUE! |  | #VALUE!
0 |  | 0 | 0 |  | 0 |  | 0 | 0 |  | 0
UNCERTAINTY OF MEASUREMENT:
The uncertainty stated is the expanded uncertainty obtained by multiplying the standard uncertainty by the coverage factor k = 2. It has been determined in accordance with the “JCGM 100:2008 Evaluation of measurement data- Guide to the Expression of Uncertainty in measurement ”. The value of the measurand lies within the assigned range of values with a probability of approximately 95%.
Page 1 of 2
STANDARD USED :
Name of Standard |  |  | Make/Model |  |  | Calibration Certificate No. |  |  | Traceability
Load Cell SN F93007 |  |  | HBM/C3H3 |  |  | 1800-00040-004 |  |  | Traceable to SI through KRISS - Korea
CALIBRATION PROCEDURE:
The testing machine under calibration was subjected to specified force values which were applied on the force-proving instrument and readings were noted.
The relevant reference document for this calibration is ISO 7500-1:2018, Metallic materials – Verification of static uniaxial testing machines.
ENVIRONMENTAL CONDITIONS:
| Ambient Temperature |  |  | : (25 ± 1) °C
| Relative Humidity |  |  | : (33 ± 5) %
REMARKS:
- | The above results were those obtained at the time of calibration and refer only to the testing machine calibrated in compression. | [9 empty] | The testing machine was found to comply with the requirements to the maximum permissible values (%) for the different relative errors of the force measuring system which characterizes a testing machine in accordance with the appropriate class as given in Table 2, Item 6 of the above method.
- | The testing machine was found to comply with the requirements to the maximum permissible values (%) for the different relative errors of the force measuring system which characterizes a testing machine in accordance with the appropriate class as given in Table 2, Item 6 of the above method.
- | No adjustment was performed on the testing machine. The user should determine suitability of the testing machine for its intended use.
- | The conversion factor of 1 kgf = 0.00980665 kN is only applicable in conversion of force units.
- | The testing machine's indicator make/model: INTERCOMP PT20 with serial number of: 0307YM23003.
- | The testing machine's indicator make/model: INTERCOMP PT20 with serial number of: 0307YM23003.
- | The testing machine was calibrated at the Force Standards Laboratory, NML ITDI of SAVVICE CORPORATION at 3rd Floor BMWC Bldg., Cagayan Valley Road, Sta. Rita, Guiginto, Bulacan.
Ahdrian Camilo C. Gernale
Science Research Specialist II
Radley F. Manalo
Senior Science Research Specialist
For the Chief, National Metrology Laboratory
MARYNESS I. SALAZAR
Head, Pressure and Force Standards Section
Page 2 of 2
Calibration Certificate No. 03-2026-FORC-0087

## Certificate 5 pts (SI)
CALIBRATION CERTIFICATE
No. 03-2026-FORC-0087
Date of Calibration |  |  | : | 2026-04-22
Calibration Item |  |  | : | Axle Weighing Scale
| Capacity |  | : | 100 tf
| Measurement Range |  | : | - -
| Resolution : |  | : | - -
| Make / Model |  | : | LS 630 W. PAD; INTERCOMP PT20
| Serial No. |  | : | 0122YL21002; 0307YM23003 (Indicator)
(Indicator)
Customer |  |  | : | SAVVICE CORPORATION
3rd Floor BMWC Bldg., Cagayan Valley Road
Sta. Rita, Guiginto, Bulacan
MEASUREMENT RESULTS:
|  | Applied Force (F std ) | Indicated Force (F ind) |  | Deviation (F ind - F std) |  | Relative Expanded Uncertainty |  | Relative Accuracy Error
|  | kgf | kgf |  | kgf |  | % |  | %
|  | 0 | 0 |  | 0 |  | - - |  | - -
|  | - - | - - |  | #VALUE! |  | - - |  | - -
|  | - - | - - |  | #VALUE! |  | - - |  | - -
|  | - - | - - |  | #VALUE! |  | - - |  | - -
|  | - - | - - |  | #VALUE! |  | - - |  | - -
|  | - - | - - |  | #VALUE! |  | - - |  | - -
UNCERTAINTY OF MEASUREMENT:
The uncertainty stated is the expanded uncertainty obtained by multiplying the standard uncertainty by the coverage factor k = 2. It has been determined in accordance with the “JCGM 100:2008 Evaluation of measurement data- Guide to the expression of uncertainty in measurement”. The value of the measurand lies within the assigned range of values with a probability of approximately 95%.
Page 1 of 2
STANDARD USED :
Name of Standard |  |  |  | Make/Model |  | Calibration Certificate No. |  |  |  | Traceability
Force Measuring Instrument SN - - |  |  |  | - - |  | 02-2025-FORC-0072 |  |  |  | Traceable to the SI through NMD-ITDI
CALIBRATION PROCEDURE:
The axle weighing scale () under calibration was subjected to specified force values which were applied on the working standard and readings were noted.
The axle weighing scale was subjected to specified force values in comparison with force reference standard values. Three (3) series of increasing force values were applied to the axle weighing scale.
The relevant references for this axle weighing scale calibration are the ISO 7500-1:2018 "Metallic materials-Calibration and verification of static uniaxial testing machine'' and TP-PFSS-F-FORC-02: (version 1) "Calibration of Axle Weighing Scales" .
ENVIRONMENTAL CONDITIONS:
| Ambient Temperature |  |  | #DIV/0!
| Relative Humidity |  |  | #DIV/0!
REMARKS:
- | The above results were those obtained at the time of calibration and refer only to the axle weighing scale calibrated in static compression mode.
- | The overall performance of this axle weighing scale was SATISFACTORY.
- | No adjustment was performed on the axle weighing scale. The user should determine suitability of the axle weighing scale for its intended use.
- | The conversion factor is 1 kgf = 9.80665 kN.
AHDRIAN CAMILO C. GERNALE
Science Research Specialist II
RADLEY F. MANALO
Senior Science Research Specialist
Issued under the autority of:
MICHAEL JASON A. SOLIS
Chief
[6 empty] | Date issued: _________________________ |  | 2026-05-19
-End of Report-
Page 2 of 2
Calibration Certificate No. 03-2026-FORC-0087

## Records
Form No. F5.5.2
Version 2 Page 1 of 1
CALIBRATION RECORD
(Use as first page of Data Sheet for any calibration or measurement work performed.)
Request Reference No.: |  |  | 03-2026-FORC-0087
Instrument: |  |  | Axle Weighing Scale
Make / Model No. |  |  | LS 630 W. PAD
Customer: |  |  | SAVVICE CORPORATION
Environmental Conditions: |  | #DIV/0! |  |  |  | #DIV/0!
Calibration Procedure: |  |  | ISO 7500/1
Date of Calibration: |  | 2018-02-21 | 2026-04-22 |  | Calibrated by: |  |  | RFM
Standards and Equipment Used:
Description |  | Make / Model No. |  | Serial No. |  |  |  |  | Date of Next Calibration
working standard (UTM) |  | - - |  | SN - -
Thermohygrometer |  | TandD/TR-73U |  | 104-104
Remarks:
Checked by:

## Others db
Types of Machine | Name | Position | Unit | equivalence to kN
Box Compression Machine | Ahdrian Camilo C. Gernale | Science Research Specialist II | kgf | 0.0098067
Axle Weighing Scale | Radley F. Manalo | Senior Science Research Specialist | lbf | 0.0044482
Universal Testing Machine | Glenn Joseph C. Franes | Science Research Specialist I | tonf | 9.8066
Edge Crush Testing Machine | Maryness I. Salazar | Section Head | n | 0.0001 | 1000 | kg | 1 | ton
Ring Crush Testing Machine | Pierre Jordan M. Mendoza | Science Research Specialist I
Compression Testing Machine
Tensile Testing Machine
Force Measuring Instrument

## 123 (hidden)
Force, N
0 | 0
196.133 | 0.19612
392.266 | 0.39222
588.399 | 0.58832
748.532 | 0.78444
980.665 | 0.98056
1176.798 | 1.1767
1569.064 | 1.3728
1765.197 | 1.569
1961.33 | 1.7651